## Dilution concentration gravi
Gravimetric preparation of calibration liquids
This sheet shall help you to calculate the correct amounts of tracer
Fill in the red marked cells with what you measure
|  | Base solution/Standard concentration (g/L) |  | 100.3
|  | Base solution/Standard density (kg/L) |  | 1.0403 | # measured
|  | Bucket volume (L) |  | 5
|  | Dilution 1 (S1) |  | Dilution 2 (S2) |  | Dilution 3 (S3) |  | CALIBRATION FLUID 1 |  | CALIBRATION FLUID 2 |  | CALIBRATION FLUID 3 |  | CALIBRATION FLUID 4 |  | CALIBRATION FLUID 5
Lab preparation | Initial solution (to be collected) | Volume (mL) | Collected mass (g) | Theoretical volume (mL) | Collected mass (g) | Theoretical volume (mL) | Collected mass (g) | Theoretical volume (mL) | Collected mass (g) | Theoretical volume (mL) | Collected mass (g) | Theoretical volume (mL) | Collected mass (g) | Theoretical volume (mL) | Collected mass (g) | Theoretical volume (mL) | Collected mass (g)
|  | 4.9697 | 5.17 | 4.99 | 4.99 | 1000 | 1000 |  | 8 |  | 8 |  | 8 |  | 10 |  | 10
| Final solution (mass in g)/size of mixing flask (ml) | 1000 |  | 1000 |  | 1000 |  | 0 | µg (theoretical mass) | 0 | µg (theoretical mass) | 0 | µg (theoretical mass) | 0 | µg (theoretical mass) | 0 | µg (theoretical mass)
| Dilution ratio | Theoretical | Actual | Theoretical | Actual | Theoretical | Actual | 20.7006 | µg (actual mass) | 20.7006 | µg (actual mass) | 20.7006 | µg (actual mass) | 25.8757 | µg (actual mass) | 25.8757 | µg (actual mass)
|  | 0.0049697 | 0.0049697 | 0.00499 | 0.00499 | 1 | 1
| Final dilution ratio | 2.4799e-05
| Concentration (mg/L) | 518.551 |  | 2.5876 |  | 2.5876
| Equipment used | 10mL pipette/syringe & 2L(1L?) volumetric flask |  | 10mL pipette/syringe & 2L(1L?) volumetric flask |  | 10mL pipette/syringe & 2L(1L?) volumetric flask
| Production | Dilution of 10 mL base solution to 2L(1L?) water |  | Dilution of 10 mL of solution S1 with 2L(1L?) water |  | Dilution of 10 mL of solution S2 with 2L(1L?) water
| Volume d'ajout (mL) |  |  | 0.5
| Volume du récipient de calibration (mL) |  |  | 500
| Taux de dilution de la calibration (pour un ajout) |  |  | 0.001
| Taux de dilution total (pour un ajout) |  |  | 2.4799e-08
| Concentration finale en µg/l (après un ajout) |  |  | 2.5876
| Matériel utilisé et réalisation |  |  | Ajout de 500µL de solution de calibration avec la µ-pipette dans le récipient isotherme de 500mL (rempli avec le ballon jaugé en plastique)
| Current concentration of 50 ml in 5L. Transforming the g value of solution into g value of actual tracer powder - (µg/L) = ppb | 5185.51 |  | 25.8757 |  | 25.8757
Actual calibration [ppb]
[6 empty] | Calibration series (final) in 5L bucket | 1 step | 4.1335
[7 empty] | 2 step | 8.2538
|  | Input values |  |  |  |  | 3 step | 12.361
[7 empty] | 4 step | 17.4766
[7 empty] | 5 step | 22.572

## Help on injection preparation
Calculations based on discharge and time of travel hypothesis:
| Hypothesis | low | high |  |  | Input values
| Discharge (m³/s) : | 0.25 | 2 |  |  |  | 0.4375
| Tracer time of travel (s) : | 200 | 1800 |  |  |  | 400
Sample to prepare (ml) : |  | Estimated discharge (m³/s)
|  | 0.25 | 0.6875 | 1.125 | 1.5625 | 2
Estimated time of travel (s) | 200 | 4.985 | 13.7089 | 22.4327 | 31.1565 | 39.8804
| 600 | 14.9551 | 41.1266 | 67.2981 | 93.4696 | 119.6411
| 1000 | 24.9252 | 68.5444 | 112.1635 | 155.7827 | 199.4018 |  | a = | 0.01
| 1400 | 34.8953 | 95.9621 | 157.0289 | 218.0957 | 279.1625 |  | b = | 0.1003
| 1800 | 44.8654 | 123.3799 | 201.8943 | 280.4088 | 358.9232 |  | a - target concentration in mg/L --> here 10 µg/L = 10 ppb
b - tracer liquid concentration [g(tracer)/g(water)]
[8 empty] | a : base value = | 0.003
m (g) = Q (m3/s) * T (s) * a / b